--------------------------------------------------------------------------

Homemade Microgram Electrobalances

by Shawn Carlson

From Scientific American June 1996

Scan and OCR by Swedish Infomania

--------------------------------------------------------------------------

Microgram balances are clever devices that can measure fantastically tiny

masses. Top-of-the-line models employ an ingenious combination of

mechanical isolation, thermal insulation and electronic wizardry to produce

repeatable measurements down to one tenth of a millionth of one gram. With

their elaborate glass enclosures and polished goldplated fixtures, these

balances look more like works of art than scientific instruments. New

models can cost more than $10,000 and often require a master's touch to

coax reliable data from background noise.

But for all their cost and outward complexity, these devices are in essence

quite simple. One common type uses a magnetic coil to provide a torque that

delicately balances a specimen at the end of a lever arm. Increasing the

electric current in the coil increases the torque. The current required to

offset the weight of the specimen is therefore a direct measure of its

mass. The coils in commercial balances ride on pivots of polished blue

sapphire. Sapphires are used because their extreme hardness (only diamonds

are harder) keeps the pivots from wearing. Sophisticated sensing devices

and circuitry control the current in the coil - which is why microgram

electrobalances are so pricey.

And that is good news for amateurs. If you are willing to substitute your

eyes for the sensors and your hands for the control circuits, you can build

a delicate electrobalance for less than $30.

George Schmermund of Vista, Calif., made this fact clear to me. For more

than 20 years, Schmermund has run a small company called Science Resources,

which buys, repairs and customizes scientific equipment. Although he may be

an austere professional to his clients, I know him to be quite the free

spirit who spends time in the business world only so he can make enough

money to indulge his true passion - amateur science.

Schmermund already owns four expensive commercial microgram balances. But

in the interest of advancing amateur science, he decided to see how well he

could do on the cheap. His ingenious ploy was to combine a cheese board and

an old galvanometer, a device that measures current. The result was an

electrobalance that can determine weights from about 10 micrograms all the

way up to 500,000 micrograms (0.5 gram).

The precision of the measurements is quite impressive. I personally

confirmed that his design can measure to 1 percent masses exceeding one

milligram. Furthermore, it can distinguish between masses in the

100-microgram range that differ by as little as two micrograms. And

calculations suggest that the instrument can measure single masses as

slight as 10 micrograms (I didn't have a weight this small to test).

The crucial component, the galvanometer, is easy to come by. These devices

are the centerpiece of most old analog electric meters�the kind that use a

needle mounted on a coil. Current flowing through the coil creates a

magnetic field that deflects the needle. Schmermund's design calls for the

needle, mounted in the vertical plane, to act as the lever arm: specimens

hang from the needle's tip.

Electronic surplus stores will probably have several analog galvanometers

on hand. A good way to judge the quality is to shake the meter gently from

side to side. If the needle stays in place, you're holding a suitable coil.

Beyond this test, a strange sense of aesthetics guides me in selecting a

good meter. It is frustratingly difficult to describe this sense, but if

I'm moved to say, "Now this is a beautiful meter!" when I look it over, I

buy it. There is a practical benefit to this aesthetic fuzziness. Finely

crafted and carefully designed meters usually house exquisite coils that

are every bit as good as the coils used in fine electrobalances, sapphire

bearings and all.

To build the balance, gently liberate the coil from the meter housing,

being careful not to damage the needle. Mount the coil on a scrap sheet of

aluminum [see illustration on opposite page]. If you can't use aluminum

sheet metal, mount the coil inside a plastic project box. To isolate the

balance from air currents, secure the entire assembly in a glass-covered

cheese board, with the aluminum sheet standing upright so that the needle

moves up and down. The two heavy guard wires cannibalized from the meter

are mounted on the aluminum support to constrain the needle's range of

motion.

Epoxy a small bolt to the aluminum support, just behind the needle's tip.

The needle should cross just in front of the bolt without touching. Cover

the bolt with a small piece of construction paper, then draw a thin

horizontal line across the center of the paper. This line defines the zero

position of the scale.

The specimen tray that hangs from the needle is merely a small frame

homefashioned by bending noninsulated wire. The exact diameter of the wire

is not critical, but keep it thin: 28-gauge wire works well. A tiny circle

of aluminum foil rests at the base of the wire frame and serves as the tray

pan. To avoid contamination with body oils, never touch the tray (or the

specimen) with your fingers; rather always use a pair of tweezers.

To energize the galvanometer coil, you'll need a circuit that supplies a

stable five volts [see illustration below]. Do not substitute an AC-to-DC

adapter for the batteries unless you are willing to add filters that can

suppress lowfrequency voltage fluctuations, which can leak into the system

from the adapter. Fluctuations as tiny as 0.1 millivolt will sharply reduce

your ability to resolve the smallest weights.

The device uses two precision, 100kilohm, 10-turn, variable resistors (also

called potentiometers or rheostats) - the first to adjust the voltage

across the coil and the second to provide a zero reference. A 20-microfarad

capacitor buffers the coil against any jerkiness in the resistors'response

and helps in making any delicate adjustments to the needle's position. To

measure the voltage across the coil, you'll need a digital voltmeter that

reads down to 0.1 millivolt. Radio Shack sells handheld versions for less

than $80. Using a five-volt power supply, Schmermund's scale can lift 150

milligrams. For larger weights, replace the type 7805 voltage-regulator

chip with a 7812 chip. It will produce a stable 12 volts and will lift

objects weighing nearly half a gram.

To calibrate the scale, you'll need a set of known microgram weights. A

single high-precision calibrated weight between one and 100 micrograms

typically costs $75, and you'll need at least two. There is, however, a

cheaper way. The Society for Amateur Scientists is making available for $10

sets of two calibrated microgram weights suitable for this project. Note

that these two weights enable you to calibrate your balance with four known

masses: zero, weight one, weight two and the sum of the two weights.

To make a measurement, begin with the scale pan empty. Cover the device

with the glass enclosure. Choke down the electric current by setting the

first resistor to its highest value. Next, adjust the second resistor until

the voltage reads as close to zero as you can set it. Write down this

voltage and don't touch this resistor again until you have finished your

entire set of measurements. Now turn up the first resistor until the needle

sinks down to the lower stop, then turn it back so that the needle returns

to the zero mark. Note the voltage reading again. Use the average of three

voltage measurements to define the zero point of the scale.

Next, increase the resistance until the needle comes to rest on the lower

wire support. Place a weight in the tray and reduce the resistance until

the armature once more obscures the line. Record the voltage. Again, repeat

the measurement three times and take the average. The difference between

these two average voltages is a direct measure of the specimen's weight.

Once you have measured the calibrated weights, plot the mass lifted against

the voltage applied. The data should fall on a straight line. The mass

corresponding to any intermediate voltage can then be read straight off the

curve.

Schmermund's balance is extremely linear above 10 milligrams. The slope of

the calibration line decreased by only 4 percent at 500 micrograms, the

smallest calibrated weight we had available. Nevertheless, I strongly

suggest that you calibrate your balance every time you use it and always

compare your specimens directly with your calibrated weights.

To receive the two calibrated weights, send $10 and a self-addressed,

stamped envelope to the Society for Amateur Scientists, 4951 D Clairemont

Square, Suite 179, San Diego, CA 92117. For more information about this

project and for suggestions for experiments, send $5 to the address above

or download them for free from the SAS Web page at

http://www.thesphere.com/SAS/ or Scientific American's area on America

Online.

--------------------------------------------------------------------------